## Equipment Breakdown
Rank | Equipment ID | Equipment | Department | No. of Calls | Total Downtime (hrs) | Avg Downtime/Call (min) | % Calls Pending | Criticality | Failure Pattern | Action Pattern | Replace?
1 | — | Electric Bed | 10 th new, 10TH NEW, 10TH NORTH, 10th new, 10th New, 10th main, 10th new, 2nd floor, 2nd main, 3RD FLOOR, 3RD Main, 3d | 142 | 133.6 | 35 | 13 | Normal | bed not working; head end stuck; remote not working | NEEDS PARTS/REPLACEMENT: not working, reconnected to replace the same; replaced th caster and solve the issue | HIGH
2 | — | Monitor | 10 TH SOUTH, 2nd main, 3rd T2, 3rd t2, 3rd x, 6th T2, 6th isolation, 8th isolation, 9th Main, Ambulance, Anasthesiology, | 55 | 60.5 | 38 | 4 | Normal | spo2 probe not working; spo2 not getting; bp not getting | NEEDS PARTS/REPLACEMENT: re[placed with new one; replaced with new probe | HIGH
3 | 10THS/EB/02 | Electric Bed | ICCU, MDICU, iccu | 3 | 1.1 | 22 | 100 | Normal | not working | NEEDS PARTS/REPLACEMENT: need to replace the hi low. checked and replaced for service; checked and need to replace spare | HIGH
4 | NEUOT/NNS/01 | Neuro Navigation | Neuro OT, Neuro OT , Neuro Ot, Neuro ot, OT, neuro ot, neuro ot , neuroot | 19 | 68.2 | 98 | 0 | Normal | to check the machine; cd checking; to assist the case | assisted the case | HIGH
5 | RD/PDR/07 | Portable Dr | RADIOTHERAPY, Radiology, radiology, x ray | 20 | 38.9 | 75 | 28 | Normal | to assist se; image not sending to pacs; kv fault | NEEDS PARTS/REPLACEMENT: replaced te plug top and now ok; informed sr engr | HIGH
6 | — | Bp Apparatus | 10th main, 2nd main, 3rd Main, 3rd T2, 3rd main, 3rd t3, 4th main, 6th North, 6th T2, 6th main, 6th north, 7th South, 7t | 51 | 23.5 | 27 | 3 | Normal | bp cuff not working properly; bp cuff not working; bp apparatus complaint | NEEDS PARTS/REPLACEMENT: replaced bulb; replaced with new cuff | HIGH
7 | — | Glucometer | 10th Main, 10th new, 2nd main, 3rd main, 3rd t2, 4THX, 4th main, 4tht2, 5th main, 6th North, 6th main, 6th north, 7th ma | 38 | 30.1 | 49 | 2 | Normal | glucometer not working; glucometer shows error; glucometer is not working sent with glu strip | NEEDS PARTS/REPLACEMENT: need to replace the meter; battery replaced | HIGH
8 | — | Nebulizer | 10 TH NEW, 1st Floor, 3rd T3, 3rd t2, 4TH T2, 5th T2, 6TH north, 6th ,main, 6th T2, 6th main, 6th new, 8th Main, 8th New | 27 | 18.9 | 54 | 25 | Normal | nebulizer is not working; nebulizer for cleaning; clean the machine | NEEDS PARTS/REPLACEMENT: checked and replaced filter; cleaned and replaced filter | HIGH
9 | RD/FLURO/01 | Fluroscopy | Radiology | 4 | 8.8 | 106 | 60 | normal | image not getting; x ray switch faulty; storage full | NEEDS PARTS/REPLACEMENT: replaced console and checked working | HIGH
10 | MB/ACL/05 | Autoclave | Microbiology, microbiology | 10 | 16.2 | 88 | 45 | normal | leak from the machine; to paint the machine; at attaining pressure gauge 120 | NEEDS PARTS/REPLACEMENT: tepm issue. informed sr: engg; checked and informed sr engg | HIGH
11 | — | Foot Switch | Main OT | 3 | 1.7 | 100 | 67 | Normal | monopolar foot switch need repair; foot switch not working; bi polar footswitch not working | checked and rectified the issue | HIGH
12 | — | Stethescope | 10th New, 10th new, 3rd , 3rd T2, 3rd main, 3rd t3, 4TH Main, 5th floor new, 5th tower 2, 6th North, 6th T2, 6th main, | 41 | 24.6 | 38 | 0 | Normal | stethescope complaint; frame broken; stethescope not working | NEEDS PARTS/REPLACEMENT: replaced with new one; frame replaced | HIGH
13 | — | Neuro Navigation | Main OT, Neuro OT, Neuro ot, Neuro ot , OT, Ot, neuro ot, neuro ot, nuero ot, ot | 17 | 49.2 | 140 | 0 | critical | cd checking; optical case; o arm with navigation | NEEDS PARTS/REPLACEMENT: camera ordered need to replace | HIGH
14 | — | Steam Inhaler | 2nd main, 3rd T2, gshdu | 3 | 0.2 | 13 | 67 | Normal | steam inhaler not working; steam inhaler is not working | NEEDS PARTS/REPLACEMENT: replaced with new one | HIGH
15 | NICU/RHW/08 | Infant Warmer | NICU, nicu | 3 | 3.8 | 46 | 60 | Normal | not switching on; temp not getting | NEEDS PARTS/REPLACEMENT: replaced power supply board and socket checked working; checked and fuse blown. replaced the fuse but the problem exist. suspended board malfunction. informed | HIGH
16 | — | Warmer | Anaesthesiology, CTOT, MOT - 3 , Main ot, NICU, anaethesiology, main OT, main ot, mot - 8 | 6 | 5.9 | 59 | 50 | Normal | warmer not working; warmer not working showing change filter; cradle cover broken. | NEEDS PARTS/REPLACEMENT: cover replaced and checked ok.; replaced the probe | HIGH
17 | — | Laryngoscope | 3rd Main, 3rd main, 6th T2, 6th south pead, CSICU, CTOT, CTOT Anasthesia, ER, ICCU, MDICU, MOT, Main OT, Main ot, Mdicu, | 27 | 31 | 62 | 0 | Normal | blade not working; bulb not working; laryngoscope not working | NEEDS PARTS/REPLACEMENT: replaced bulb; bulb replaced | HIGH
18 | — | Fridge Thermometer | 2nd main, 3T 2, 3rd main, 4th T2, 5th extension, 6th North, 6th north, 6th south, 6th south paediatric ward, 7th main, 7 | 34 | 20.6 | 36 | 3 | Normal | fridge thermometer not working; digital thermometer not working; fridge thermometer due date was over | NEEDS PARTS/REPLACEMENT: battery replaced; checked and replaced battery | HIGH
19 | — | Ventilator | 8th deluxe, 8th isolation, AMBULSNCE, Ambulance , MDICU, Mdicu, Mmdicu, NICU, NSICU, Neuro Hdu, Neuro ICU, Onco ICU, RTI | 21 | 34 | 57 | 3 | Normal | demo; new installation; to set the ventilator | NEEDS PARTS/REPLACEMENT: need to service; replaced with 2 nos flow sensor | HIGH
20 | MICU/VEN/02 | Ventilator | 8th deluxe, 8th dlx, 8th isolation, 8th main, GMICU, MDICU, Onco icu, THDU, mdicu | 8 | 6.7 | 34 | 42 | Critical | device failure; neb button not working; flow sensor inop | NEEDS PARTS/REPLACEMENT: replaced flow sensor and done calibration; replaced the knob and working ok | HIGH
21 | — | Infusion Pump | 10 th main, 3rd extn, 3rd main, 3rd t2, 3rd x, 4TH x, 4th main, 4th t2, 4thx, 6th floor, 6th main, 6th old, 7th main, 8T | 24 | 20 | 41 | 9 | Normal | to set the pump; showing alarm; alarm | NEEDS PARTS/REPLACEMENT: recommended to replace with another one; asked to replace the pump. taken for service | HIGH
22 | — | Syringe Pump | 6th floor, 6th isolation, 6th north, 8th South, 8th isolation, In house, In house , LTICU, MDICU, MOT, Ortho OT, TICU, m | 22 | 21.3 | 64 | 9 | Normal | syringe pump is nt working; syringe off; syringe pump no: 6 and 9 for calibration | NEEDS PARTS/REPLACEMENT: need to change battery; replaced the screw | HIGH
23 | OROP/LGT/03 | Spot Light | Foot and ankle, OROP | 3 | 2.3 | 35 | 50 | Normal | wheel base broken; wheel found damage; caster wheel open | NEEDS PARTS/REPLACEMENT: checked and need to replace | HIGH
24 | MDICU/BGA/01 | Abg Machine | HDICU, MDICU, Mdicu, mdicu | 18 | 17.6 | 36 | 17 | Normal | to do cvp; cvp error; hct not getting | NEEDS PARTS/REPLACEMENT: repolaced with new catridge; inserted new catridge and found ok | HIGH
25 | — | Trolly | 6th Floor, 9th main , Emergency, GSHDU, Microbiology, Old Emergency , Recovery, emergency, er | 10 | 7.6 | 51 | 36 | Normal | to replace the wheel and hydrolic pump; trolly damaged; wheel stuck | NEEDS PARTS/REPLACEMENT: replaced the wheel and hydrolic pump; need replacement | HIGH
26 | — | Ecg Machine | 10TH main, 10th main, 3rd main, 5TH T2, 5th extensition, 5th old, 6th main, 8TH Main, 8th main, Ambulance, Biomedical, E | 19 | 21.9 | 53 | 7 | Normal | installation; ecg machine not working; ecg battery low | NEEDS PARTS/REPLACEMENT: replaced with new one; replaced chest lead and bulb | HIGH
27 | BB/TACE/01 | Component Extractor | Blood Bank, Blood bank, blood bank | 5 | 5.7 | 49 | 38 | Normal | breakdown; door not getting closed; rbc entering plasma | NEEDS PARTS/REPLACEMENT: sealing not happening. informed sr engg; informed sr engg | HIGH
28 | OT/ICG/01 | Icg System | Main OT, OT 3, main ot | 4 | 4.6 | 55 | 40 | Normal | icg not working; icg problem; fog in image | NEEDS PARTS/REPLACEMENT: checked with eng. need to replace, light cable and bulb | HIGH
29 | — | Microscope | Histopathology, Microbiology, heamatology, hematology, histopathology | 3 | 12 | 120 | 33 | normal | high power not working; issue with the power cord; demo | NEEDS PARTS/REPLACEMENT: indented for new one; need to replace board | HIGH
30 | — | Diathermy | Main OT, Main ot, main ot | 7 | 11.2 | 112 | 29 | Normal | monopolar pedals cutting is not working; bi polar port not working; monopolar coagulation switch is keep on pressed, while pressing | taken for service | HIGH
31 | — | Biosafety Cabinet | 5th main, 7th floor new block, 7th new, 7th oncology, microbiology | 3 | 4.5 | 54 | 40 | Normal | installation; new installation duct issues; lamp not working | assisted the se | HIGH
32 | NNS/01 | Neuro Navigation System | Neuro , Neuro OT, Neuro ot, Neuro ot , OT, neuro or, neuro ot, neuro ot | 14 | 29.8 | 99 | 0 | Normal | cd checking; to check the machine after case; to set the machine | assisted the case | HIGH
33 | — | Surgical Diathermy | Head and neck, Main OT, OT | 3 | 9.9 | 198 | 33 | Normal | installation | checked and rectified | HIGH
34 | RD/PDR/08 | Portable X Ray | NSICU, Radiology, Radiology , SICU, radiology | 9 | 16.9 | 56 | 17 | Normal | kv fault; detector connection failed; image not sending to pacs | NEEDS PARTS/REPLACEMENT: replaced the plug top and found ok; reconnected the collimeter connection and need to check | HIGH
35 | — | Pneumatic System | 2nd machine, 6th new, 7th south, Biochemistry, Chemo ward, Laboratry, Onco Pharmacy, Onco pharmacy, Pneumatic, Pneumatic | 21 | 20.1 | 53 | 0 | Normal | sample missing; keypad issue; smaple missing | NEEDS PARTS/REPLACEMENT: replaced the keypad. now working good; replaced sensor with se and now ok | HIGH
36 | — | Cusa | Neuro OT, OT, OT 2, OT 8 , Robotic OT, ot 7 | 5 | 4.8 | 48 | 33 | Normal | handpiece broken; to assist the case; not working properly | NEEDS PARTS/REPLACEMENT: checked, pin broken need to replace | HIGH
37 | — | Biosafty Cabinet | 7th new, 7th south, BMTU, Onco ICU/7th N, microbiology, oncology | 5 | 4.1 | 61 | 33 | normal | alarm sound; to check the installation site; installation | NEEDS PARTS/REPLACEMENT: flow sensor need to tight; checked with se need to replace | HIGH
38 | — | Abg Machine | Biochemistry, Biohcemistry, CSICU, Csicu, LTICU, MDICU, biochemistry, csicu, mdicu, sicu | 13 | 15.4 | 62 | 12 | normal | for catridge replacement; catridge expired; for performing cvp | NEEDS PARTS/REPLACEMENT: checked and replaced the catridge; checked and found ok,informed sr engr | HIGH
39 | RD/MRI/01 | Mri | MRI, Radiology, radiology | 8 | 10.8 | 72 | 22 | Normal | mri chiller issue; cd writer not working; to assist se in service | NEEDS PARTS/REPLACEMENT: replaced the bushing filter | HIGH
40 | RD/MAM/02 | Mamography | Mamography, Radiology, radiology | 11 | 20.8 | 104 | 8 | Normal | work station switching on and off; calibration due alarm; mammo unit error 12.41.0.0 | NEEDS PARTS/REPLACEMENT: checked found x ray tube error need to replace; tube replaced | HIGH
41 | OT/CARM/02 | Carm | OT, main ot, ot | 3 | 9.2 | 79 | 29 | Normal | c arm wheel stuck; image clarity issue; wheel dmaaged | NEEDS PARTS/REPLACEMENT: need to replace wheel | HIGH
42 | — | Incubator | CSSD, NICU, cssd, nicu | 4 | 4.3 | 33 | 33 | normal | temp probe not working; temp checking; probe insulation damaged | NEEDS PARTS/REPLACEMENT: replaced the cable; informed company for replacement | HIGH
43 | — | Steamer | 3rd T2, 4th extension, 5th Main, 6th Main, 8th north, NHDU, TICU, bmt | 8 | 6.9 | 69 | 25 | Normal | steamer not working; console note working; warmer not working | NEEDS PARTS/REPLACEMENT: coil burned and need to contempnation | HIGH
44 | — | Ot Table | CTOT, MAIN OT, Main OT, Neuro ot, OT 3, OT 4, OT 6, OT 7, OT4, Ortho ot, Plastic surgery, Robotic OT, endoscopy, neuro o | 14 | 19.5 | 65 | 5 | Normal | remote not working; tiut not working; table switch complaint | NEEDS PARTS/REPLACEMENT: checked and replaced the remote | HIGH
45 | SICU/MON/22 | Monitor | SICU, In house, SICU | 7 | 3.5 | 35 | 29 | Non critical | cvp not monitoring; to change date and time; to remove the monitor from the stand | NEEDS PARTS/REPLACEMENT: replaced the power supply and found ok | HIGH
46 | RD/US/12 | Ultrasound Machine | Radiology, Radiology , Ultrasound , radiology, ultrasound | 6 | 15.9 | 80 | 17 | Normal | printer not working; pink line in the touch panel; checked worklist not coming | NEEDS PARTS/REPLACEMENT: replaced the jam paper and now working ok; checked and need replaced the touch panel | HIGH
47 | CSSD/SS/05 | Steam Sterilizer | CSSD, cssd | 3 | 7.5 | 64 | 29 | Normal | stainer inside the chamber is damaged; ups intruptted; stainer for chamber complaint | NEEDS PARTS/REPLACEMENT: replaced the strainer, machine is working fine; condensor replaced amd checked working | HIGH
48 | CATH/MON/02 | Monitor | Cah lab, Cath lab, cathlab | 3 | 2.9 | 58 | 33 | Normal | cardiac monitor not working; not getting bp; not switching on | taken for service | HIGH
49 | EMR/US/02 | Ultrasound Machine | Emergency, Radiology, radiology | 4 | 19.5 | 167 | 14 | Normal | linear probe not working; echo probe issue; echo probe has low clarity | NEEDS PARTS/REPLACEMENT: informed sr engg; assisted se in checking the probe . need to replace new one | HIGH
50 | ENT/AUD/02 | Audiometer | Audiology, ENT | 3 | 2.2 | 43 | 33 | Normal | to assist with se; audiometer cable complaint; push button not working | assistedv se | HIGH
51 | SICU/MON/16 | Monitor | 2nd Floor , OSHDU, Onco HDU, onco HDU, onco hdu, onco sx hdu, oshdu | 4 | 4.6 | 46 | 29 | Non critical | mmx not detecting; bp not getting; mms not getting | NEEDS PARTS/REPLACEMENT: reconnected mmx board connection but need to replace speaker and power supply board; need to ckeck after 2 days | HIGH
52 | — | Electric Cot | 10 th north, 10th South, 3rd main, 4th main, 4th old, 8th isolation, day care | 6 | 2.2 | 22 | 29 | Normal | head end not working; full up not working | NEEDS PARTS/REPLACEMENT: checked and need to replace head motor; checked and need to take for service` | HIGH
53 | 10THS/EB/06 | Electric Bed | 10th New, 10th new, 9th new | 3 | 1.6 | 31 | 33 | Normal | bed not working; head end not working | changed the power supply board and working fine | HIGH
54 | BB/AIS/01 | Automated Immunoassay System | Blood Bank, Blood bank | 3 | 1.6 | 32 | 33 | Normal | hiv kit not ready; automatically turning off; no battery backup | NEEDS PARTS/REPLACEMENT: call logged and informed sr eng | HIGH
55 | TICU/US/01 | Portable Ultrasound M/C | TICU, Ticu | 3 | 1.6 | 32 | 33 | Normal | to assist se in pm; m/c turn off; to rearrange and stow the cable | NEEDS PARTS/REPLACEMENT: checked and rectified. need to replace the battery | HIGH
56 | HP/GWS/01 | Grosing Station | Histopathology | 3 | 1.5 | 30 | 33 | Normal | issue with formalin monitor and tap issue; pm done; drain switch not working | checked but problem exist | MEDIUM
57 | LTICU/EB/04 | Electric Bed | 6th isolation, 9th New, Neuro HDU | 3 | 1.5 | 29 | 33 | Normal | not working | NEEDS PARTS/REPLACEMENT: replaced for service | MEDIUM
58 | BB/PAI/02 | Platelet Incubator | Blood Bank, Blood bank, blood bank | 3 | 1.3 | 26 | 33 | Normal | to assist se in service; temp high; display blinking | assisted se in service . rectified the issue | MEDIUM
59 | — | Head Light | CTOT, ENT OPD, Main OT | 3 | 1.2 | 25 | 33 | Normal | lens cover loose; head light complaint; new installation | NEEDS PARTS/REPLACEMENT: need to taken for service to company | MEDIUM
60 | 7THNEW/EB/10 | Electric Bed | 4TH Main, 4th Old, 6th main | 3 | 0.9 | 18 | 33 | Normal | electric bed not working | NEEDS PARTS/REPLACEMENT: need to change control box; need to service | MEDIUM
61 | NEPH/DM/15 | Dialysis Machine | Dailysis, Dialysis, dialysis | 3 | 0.9 | 27 | 33 | Normal | pm due; coupling leakage; water leak | pm done | MEDIUM
62 | — | Bp Cuff | 10TH New, 10th main, 3rd main, 4th new, 9th main, Anesthesia, Chemoward, Dialysis, Emergency, LTICU, Main OT, NSICU, Ort | 22 | 10.2 | 25 | 0 | Normal | bp cuff complaint; bp cuff not working; damaged cuff | NEEDS PARTS/REPLACEMENT: replaced with new one; replaced cuff and bulb | MEDIUM
63 | OT/AW/15 | Anesthesia Machine | OT, OT 8, Robotic OT, ot , ot 8 | 5 | 12.1 | 121 | 17 | Normal | values not getting; leak test fail; device failure | NEEDS PARTS/REPLACEMENT: replaced the tubings and water tap, now leak test passed; standby pixton replaced and checked working | MEDIUM
64 | — | Ultrasound Machine | 3RD Executive, 3rd executive, 5th new, ECHO, LTICU, Radilogy, Radiology, ULTRASOUND, Ultrasound, Ultrasound Room, csicu, | 12 | 14.1 | 29 | 6 | Normal | image not sending to pacs; not sending to pacs; images not sending to pacs | NEEDS PARTS/REPLACEMENT: not ok need to apply the image to cd; checked and temperorly rectified,informed sr.engr | MEDIUM
65 | — | Pnuematic System | 6th North, biochemistry, pnuematic | 4 | 4.5 | 68 | 25 | normal | line 2 error; shooter system not working; error | checked and done | MEDIUM
66 | — | Feeding Pump | 4th main, 6th main, 6th south, 8TH EX, 8th south, 9th south , GSHDU, sicu | 5 | 3.2 | 32 | 25 | Normal | power cord not working feeding pump; the feeding pump is alarming | NEEDS PARTS/REPLACEMENT: checked but issue exists. temporarily arranged for new pump. taken for service | MEDIUM
67 | RD/MRI/02 | Mri | MRI, MRI 02, Radiology, radiology | 6 | 4.9 | 33 | 22 | Normal | communication error; printer not working; showing table error | checked and rectified the isssue | MEDIUM
68 | — | Neuro Navigation System | Neuro OT, Neuro OT , Neuro ot | 7 | 24.5 | 210 | 0 | Normal | case assisted; o arm case; forcase assisit | assisted the case | MEDIUM
69 | CSICU/BGA/01 | Abg Machine | CSICU, CSICU , Csicu, csicu | 13 | 12.2 | 43 | 6 | Normal | to do cvp; to doing cvp; to insert catidge | NEEDS PARTS/REPLACEMENT: loaded the catridge with barcode new one; machine is now working okay. need to call se for further examination of error | MEDIUM
70 | OT/CO2/02 | Co2 Laser | ENT, ENT OP, OT, ot | 6 | 9.2 | 92 | 17 | critical | to shift the machine; to shift the machine to ent; to shft the m/c | NEEDS PARTS/REPLACEMENT: need to replace the same; checked and working ok. ifproblem exist need to assist case | HIGH
71 | RD/US/03 | Ultrasound Machine | 3rd Executive, 3rd executive, Radiology, radiology, ultrasound executive | 5 | 5.2 | 35 | 22 | Normal | image is not sending to pacs; not sending to pacs; print not working | NEEDS PARTS/REPLACEMENT: checked and informed sr.engr | MEDIUM
72 | — | Dozee Vital Sign Monitoring | 10th main, 10th new, 3rd floor, 3rd main, 3rd,4th floor, 4th Floor , 4th floor, 4th floor (414), 4th main, 6th floor, 6t | 9 | 21.7 | 48 | 0 | Normal | for installation; for checking electrical sockets; installation of new sockets | NEEDS PARTS/REPLACEMENT: checked and informed electrical department. need to resume installation after electrical works; installed new sockets in 2 rooms of 3rd floor | MEDIUM
73 | RD/US/10 | Ultrasound Machine | 3rd excecutive, 3rd exe, Radiology, Ultrasound executive, radiology | 9 | 7.1 | 33 | 15 | Normal | thyroid probe stuck; to assist se in checking the machine; printer not working | NEEDS PARTS/REPLACEMENT: informed sr engr | MEDIUM
74 | — | Trolley | 8th south, Emergency, emergency | 4 | 2.7 | 41 | 25 | Normal | lock not working; gas spring replace; pedal broken | NEEDS PARTS/REPLACEMENT: checked and found wheel complaint, need to replace; replaced the gas spring and now ok | MEDIUM
75 | — | Portable Dr | Radiology, csicu, onco icu, radiology | 10 | 7.5 | 30 | 13 | Normal | image not sending to pacs; detector connection fail; showing error | NEEDS PARTS/REPLACEMENT: checked and replaced the plug top | MEDIUM
76 | — | C Arm | Main OT, OT, Radiology, ortho ot, radiology | 7 | 10.7 | 80 | 12 | Normal | installation; qa; c arm training | installed and checked working | MEDIUM
77 | — | Thermometer | 6th main, 6th north, 8th deluxe, All labs, CTOT, GF Pharmacy, HLA, IP Pharmacy, OT pharmacy, Pharmacy, Picu, Radiology, | 20 | 7.6 | 22 | 0 | Normal | fridge thermometer not working; damaged; calibration of thermometer | NEEDS PARTS/REPLACEMENT: thermometer replaced; replaced with new one | MEDIUM
78 | — | Ultrasound | Gynec OPD, Radiology, Radology, Ultrasound executive, radiology | 5 | 11.7 | 88 | 12 | Normal | pacs not available; hang issue; probe not getting select | NEEDS PARTS/REPLACEMENT: checked by sr.eng,need to replace acquisition board | MEDIUM
79 | NICU/INC/01 | Incubator | NICU, nicu | 5 | 5.9 | 59 | 17 | Normal | temp issue; fan error; high temperature | NEEDS PARTS/REPLACEMENT: temp board replaced; checked temp manually and checked informed sr.engr | MEDIUM
80 | — | Weighing Machine | 10th new, 10th floor , 3RD EXE, 3rd floor, 5th New, 6th Isolation, 6th isolation, 7TH floor, 8TH Floor , 8th new , 9th | 9 | 8.9 | 27 | 9 | Normal | to check the equipment; change powecode; to set the machine | checked the equipment with se | MEDIUM
81 | RD/P-US/01 | Portable Ultrasound | 8th north, Radiology, Radiology | 5 | 2.3 | 28 | 20 | Normal | no probe is detected; assist se in checking the probe; display issue | checked and reinserted the probe | MEDIUM
82 | NICU/VEN/03 | Ventilator | 8th isolation, NICU, Nicu, nicu | 9 | 16.7 | 42 | 0 | Critical | leak; to check the ventilator; to check the machine | NEEDS PARTS/REPLACEMENT: reconnected the property and informed sr eng; checked found o2cell need to replace | HIGH
83 | CSSD/PSR/02 | Plasma Sterilizer | CSSD, Cssd, cssd | 6 | 8.4 | 63 | 12 | Normal | password and username incorrect; error showing; printer and touch not working | checked and rectified the issue | MEDIUM
84 | — | Spo2 Probe | Emergency, GMICU, GSHDU, LTICU, Labour room, MDICU, NSICU, Onco icu, SDICU, SICU, emergency, iccu, mdicu, pathology, sdi | 18 | 6.9 | 23 | 0 | Normal | spo2 probe not working; probe not working properly; probe not working | NEEDS PARTS/REPLACEMENT: replaced with new one; replaced with new probe | MEDIUM
85 | SICU/VEN/02 | Ventilator | Onco ICU, SICU, SICU | 4 | 2.2 | 26 | 20 | Critical | tidal volume not getting; internal battery failure; fio2 low | NEEDS PARTS/REPLACEMENT: need to calibrate o2 sensor; replaced the battery and done checkout | HIGH
86 | — | Pulse Oximeter | 2nd Main, 4TH T2, 4th main, Main OT, OT | 5 | 1.1 | 22 | 20 | Normal | pulse oximeter not working; replaced to bionmedical department | NEEDS PARTS/REPLACEMENT: replaced probe and cable; replaced with new one | MEDIUM
87 | — | Patient Trolley | 7th main, Casuality, ENDOSCOPY, Emergency, daycare, emergency, ultrasound | 10 | 6.2 | 31 | 8 | Normal | installation; new installation; wheel stuck | installed and checked working | MEDIUM
88 | — | Pendant | Clicu, GMICU, GSHDU, ICCU, ICU's, MDICU, Main OT, Onco ICU, main OT, neuro ot | 9 | 14.6 | 73 | 0 | Normal | pm; suction problem; to do calibration with se | NEEDS PARTS/REPLACEMENT: need to service the pendant informed to eng; reconnected the phase connector and now ok. informed sr eng to reach | MEDIUM
89 | RD/US/11 | Ultrasound Machine | 3RD executive, Radiology, Ultrasound, Ultrasound , Ultrasound executive | 5 | 5.3 | 46 | 14 | Normal | pacs issue; to change the tv probe; not getting workist | NEEDS PARTS/REPLACEMENT: informed sr eng | MEDIUM
90 | RD/US/09 | Ultrasound Machine | Radiology, ULTRA SOUND, Ultrasound, Ultrasound , radiology | 8 | 8.7 | 40 | 7 | Normal | worklist not getting; artifacts; to shift machine | NEEDS PARTS/REPLACEMENT: checked and informed sr.engr | MEDIUM
91 | MICU/VEN/01 | Ventilator | 8th delx, NEURO HDU, Neuro HDU, nhdu, nshdu | 3 | 7.2 | 61 | 14 | CRITICAL | not getting on; flow sensor issue; error 401010 | NEEDS PARTS/REPLACEMENT: restated and now ok informed sr eng to recheck; replaced power suppy and rectified | HIGH
92 | NEU/NNS/01 | Neuro Navigation | Neuro OT | 5 | 18.3 | 220 | 0 | Normal | neuro oarm case; neuro case o arm; o arm case with neuro navigation | case done | MEDIUM
93 | — | Sterilizer | CTOT, Main OT, Neuro OT, Radiology | 4 | 2.8 | 28 | 17 | Normal | heating issue | NEEDS PARTS/REPLACEMENT: checked found issue, need to replace the same; serviced temperorly, need to purchace coil | MEDIUM
94 | — | Mayfield | Neuro OT, Neuro ot, OT, OT Neuro, neuro ot | 6 | 2.9 | 25 | 14 | Normal | screw tight; cant adjust the mayfield; clamping issue | lubricated and now working ok | MEDIUM
95 | SICU/VEN/01 | Ventilator | 6th isolation, SICU, SICU , Sicu, onco icu | 3 | 3.1 | 31 | 17 | normal | tidal volume not achieving; temp issue; tidal volume low | NEEDS PARTS/REPLACEMENT: informed sr eng need to check | MEDIUM
96 | LR/IFP/05 | Infusion Pump | 3rd Main, 3rd T2, 4th extension, 6th main, 8th south, MDICU | 7 | 1.6 | 16 | 14 | Normal | to set the infusion; showing alarm; alarm | checked and rectified | MEDIUM
97 | RD/CT/03 | Ct | Radiology, radiology | 5 | 3.7 | 56 | 14 | Normal | not able to plan; unwanted sound from gantry; machine stuck | checked and no issue found | MEDIUM
98 | — | Central Monitor | LTICU, NSICU, Neuro icu, SICU, main ot | 7 | 5.2 | 34 | 10 | Normal | not turn on; central monitor not working; not turning on | NEEDS PARTS/REPLACEMENT: checked and replaced with new monitor | MEDIUM
99 | — | Hfnc | GMICU, Onco ICU, PICU, lticu, onco icu | 6 | 2.3 | 20 | 14 | Normal | plug top not working; abnormal sound; fio2 not | NEEDS PARTS/REPLACEMENT: changed with new plug top (6a) | MEDIUM
100 | CSSD/SS/04 | Steam Sterilizer | CSSD, cssd | 3 | 7.2 | 54 | 12 | Normal | machine aborted; feed tank is not filling; showing alarm | NEEDS PARTS/REPLACEMENT: replaced level sensor and checked working; pm kit replaced and checked working condition | MEDIUM
101 | BC/BGA/04 | Abg Machine | Biochemistry, biochemistry | 5 | 6 | 40 | 11 | Normal | to replace the catridge; hct error; catridge change and cvp | NEEDS PARTS/REPLACEMENT: replaced the cartridge with new one; replaced catridge | MEDIUM
102 | TICU/VEN/01 | Ventilator | GMICU, NEURO HDU, ONCO ICU, Onco ICU, Renal Transplant ICU, TICU 1 , Trauma HDU, neuro hdu, nsicu, ticu | 7 | 7.5 | 35 | 7 | Critical | flow sensor inop; to set the ventilator; values not getting | NEEDS PARTS/REPLACEMENT: replaced flow sensor,resolved issue; checked and replaced the flow sensor | HIGH
103 | RD/US/07 | Ultrasound Machine | CT, Endoscopy , Radiology, Ultrasound, Ultrasound room, b2, radiology | 7 | 13.5 | 90 | 0 | Normal | to shift the machine; images not sending to pacs; image cannot be sent to pacs | NEEDS PARTS/REPLACEMENT: replaced with new one and found ok | MEDIUM
104 | BC/BGA/02 | Abg Machine | LTICU, biochemistry, lticu | 4 | 5.4 | 41 | 11 | Normal | to load catridge; to carry out cvp; old print coming out from printer | NEEDS PARTS/REPLACEMENT: chip removed and need to load new catridge | MEDIUM
105 | — | Anesthesia Machine | Brachytherapy, Cath lab, Main OT, OT, OT 7, Radiology, b2, neuro ot, ortho ot | 11 | 8.2 | 45 | 0 | Normal | leak test not pass; pm; checkout not passing | NEEDS PARTS/REPLACEMENT: replaced the wheel | MEDIUM
106 | — | Ambal Software | Endoscopy, Endoscopy , dr shawn opd, endoscopy, endoscopy | 9 | 8.3 | 45 | 0 | Normal | image not getting; image not displaying; not getting output | checked and rectified the issue | LOW
107 | — | Dozee Vital Monitoring System | 10th Main, 10th main, 3rd main, 4th main, 4thmain, 6th Main, 6th main, 7th main, 8th south, 9th floor, 9th main | 9 | 7.3 | 37 | 0 | normal | demo; training of system; set up the system for using | gave demo of system | LOW
108 | RD/AW/01 | Anesthesia Machine | Anaesthesia, MRI, Radiology, radiology | 10 | 5.8 | 32 | 0 | normal | shifting; leak test fail; leak test fail while connecting | NEEDS PARTS/REPLACEMENT: replaced o ring, cross checked by se; flow meter knob and balls is working. digital meter need to replace, informed engineer and now ok | LOW
109 | NSICU/MON/11 | Monitor | NSICU, Neuro ICU, Neuro icu, nsicu | 7 | 8.4 | 72 | 0 | Normal | arterial bp not getting; arterial not getting; ecg not getting | NEEDS PARTS/REPLACEMENT: ecg leads replaced and checked working | LOW
110 | RD/US/06 | Ultrasound Machine | Radiology, Radiology , USG Room, Ultrasound, Ultrasound 3rd, radiology | 8 | 6.9 | 35 | 0 | Normal | worklist not getting; c52e probe not working; track ball not working | NEEDS PARTS/REPLACEMENT: need to replace the probe | LOW
111 | NEUROOT/NNS/01 | Neuro Navigation | neuro ot | 4 | 10.7 | 161 | 0 | normal | to assist the case; to assist the em case | done with same | LOW
112 | OT/MON/02 | Monitor | Main OT, OT | 4 | 9.8 | 146 | 0 | Normal | etco2 not getting; ecg not getting; monitor charging to pendent | NEEDS PARTS/REPLACEMENT: checked and ok. need to change the cable | LOW
113 | CARD/TMT/04 | Tmt M/C | TMT | 3 | 10.6 | 212 | 0 | Normal | tmt belt not working; reclination is not occurred in 3rd stage; incline not incresing after stage 2 | checked and restarted then found ok | LOW
114 | NEPH/DM/17 | Dialysis Machine | Dialysis, NEPHROLOGY , dialysis | 3 | 9.5 | 190 | 0 | normal | flow alarm; coupling leak | NEEDS PARTS/REPLACEMENT: replaced blue and red coupling; replaced the degassing pump with new one | LOW
115 | — | Doro | Neuro OT, Neuro ot, neuro ot | 3 | 9.2 | 185 | 0 | Normal | doro head attachment c clip damaged; checking after service; not able to fix properly | NEEDS PARTS/REPLACEMENT: new c clip installed and now ok | LOW
116 | — | Digital Thermometer | op pharmacy, Histopathology, Medical store, Microbiology, ip pharmacy, microbiology | 8 | 3.7 | 28 | 0 | Normal | digital thermometer not working | NEEDS PARTS/REPLACEMENT: replaced the battery; battery replaced | LOW
117 | RD/DINJ/06 | Dye Injector | CT , Radiology, radiology | 3 | 9 | 270 | 0 | Normal | pm due | NEEDS PARTS/REPLACEMENT: need to replace | LOW
118 | — | Ecg Cable | ICCU, MDICU, OT 4, SDICU, emergency, gmicu, iccu | 7 | 4.6 | 46 | 0 | normal | cable broken; ecg cable not working; break down | NEEDS PARTS/REPLACEMENT: replaced with new cable; replaced with new one | LOW
119 | — | Defibrillator | 6th floor, 7th main, Ambulance, Ambulance , ER, ICCU, LTICU, Ot7, oshdu, sicu | 7 | 4.5 | 27 | 0 | Normal | installation; showing therapy delivery test fail; demo | NEEDS PARTS/REPLACEMENT: ecg leads replaced (power code not in stock returned to iccu) | LOW
120 | — | Pneumatic | 6th North , Bmt, Chemoward, IP Pharmacy, Pneumatic, biochemistry, iccu, pneumatic , pneumatic system | 4 | 7.6 | 51 | 0 | Normal | sample missing; checking the shooter; sample stuck | informed se | LOW
121 | CATHLAB/04 | Cath Lab | Cath Lab, Ctahlab, cathlab | 3 | 8.2 | 165 | 0 | Normal | mic not working; cathlab pm due; acqusition error | NEEDS PARTS/REPLACEMENT: checked and power adapter complaint. need to replace; spare replaced and working fine | LOW
122 | — | Dozee | 3rd main, 4th floor, 6th floor, 9th N, 9th new, 9th south | 7 | 3.3 | 28 | 0 | Normal | to connect dozee to new patient; dozee installation; tto check the working of equipment | NEEDS PARTS/REPLACEMENT: checked and replaced the sheet. rectified the issue | LOW
123 | LTICU/BGA/01 | Abg Machine | CTICU, LTICU, OT, lticu | 5 | 5.4 | 54 | 0 | Normal | to insert catridge; lactate error; to change cartridge and do cvp | NEEDS PARTS/REPLACEMENT: inserted new catridge. found ok; replaced the catridge and found okay | LOW
124 | OT/CO2L/02 | Co2 Laser | 0t, ENT OPD, Main ot , OT, ent opd, main ot, ot, radiology | 4 | 6.3 | 29 | 0 | normal | shifting; to shift the machine; checking the machine | shifted the machine | LOW
125 | — | Dozee Vitals Monitor | 3rd floor, 3rd,4th,6th, 4TH FLOOR, 6th floor | 4 | 6.2 | 125 | 0 | normal | for fixing socket with electrical department; for checking the availibility of rooms to do electrical work | fixed and powerwork also completed and checked the connectivity | LOW
126 | CSICU/MON/04 | Monitor | 4THF, CSICU, Csicu, csicu | 7 | 2.8 | 19 | 0 | Normal | to fix the monitor; to fix the monitor on stand; to fix on stand | fixed the monitor | LOW
127 | OT/SAR/02 | Surgical Aspirator | Main ot, OT, OT 7, OT7, ot, ot 7 | 3 | 7.1 | 71 | 0 | Critical | cusa machine setting; to assist with se; hand piece received | NEEDS PARTS/REPLACEMENT: need to change suction filter temporarly adjusted the suction tube | LOW
128 | — | Usg Machine | MDICU, Radiology, radiology | 5 | 4.6 | 46 | 0 | Normal | not turning on; flickering on bottom touch screen; probe damage | informed se | LOW
129 | CARD/ECG/04 | Ecg Machine | 3rd enew echo, 3rd excecutive, 3rd executive, 3rd extensition, Cardiology | 6 | 3.5 | 35 | 0 | Normal | print stuck; not working properly; paper jam | checked and rectified the issue | LOW
130 | NEPH/DM/14 | Dialysis Machine | Dialysis, dialysis | 6 | 3.3 | 39 | 0 | Normal | no display; blood leak; pm to be done | pm done | LOW
131 | SICU/VEN/07 | Ventilator | GMICU, NSICU, SICU, onco icu | 3 | 6.4 | 96 | 0 | normal | o2 fail; no battery backup; to set the ventilator | connected properly found working ok | LOW
132 | — | Steam Sterilizer | CSSD, ONCO, cssd | 4 | 5.3 | 46 | 0 | Normal | door not closing; door can't locked; error | NEEDS PARTS/REPLACEMENT: replaced with newone | LOW
133 | EMR/DFB/02 | Defibrillator | Emergency, Emergency , emergency, er | 6 | 3.1 | 27 | 0 | Normal | operational checkup fail; battery fail; operational checkup failed | NEEDS PARTS/REPLACEMENT: replace the battery and found ok | LOW
134 | — | Bp Apparatus, Stethescope | Blood bank, Dialysis, dialysis | 3 | 6.2 | 124 | 0 | Normal | bulb complaint, ear piece damaged; bulb damaged, frame complaint | NEEDS PARTS/REPLACEMENT: replaced bp cuff; bulb replaced, frame replaced | LOW
135 | — | Dialysis Machine | Dailysis, Dialysis, dialysis | 4 | 5.1 | 61 | 0 | Normal | concentration port leak; utilizaition; showing coupler error | checked and informed se | LOW
136 | — | Ct | CT, Radiology, radiology | 3 | 5.6 | 84 | 0 | Normal | unwanted sound; workstation issue; table cover issue | NEEDS PARTS/REPLACEMENT: replaced the tube. found working | LOW
137 | OROT/AW/01 | Anesthesia Machine | 0t 9, Orot, Ot 9, orot, ortho ot, ot | 5 | 2.8 | 28 | 0 | Critical | leakage failed; tidal volume not getting; leak test not pass | NEEDS PARTS/REPLACEMENT: flow sensor replaced and flow sensor acalibration done | LOW
138 | — | Patient Monitor | CSICU, Emergency, Neuro hdu, csicu, onco icu, sicu | 6 | 1.6 | 16 | 0 | Normal | artifacts in the ecg; to fix the monitor in the stand; not getting on | NEEDS PARTS/REPLACEMENT: checked , need to replace smps board | LOW
139 | SICU/FP/03 | Feeding Pump | 10 Th Main, 4th main, 6th Main, GSHDU, SICU | 6 | 1.4 | 17 | 0 | Normal | charging issue; not working properly; door not closing | NEEDS PARTS/REPLACEMENT: need to change the charger; adjusted the door and closed, need to service | LOW
140 | OT/AW/07 | Anesthesia Machine | cathlab, cathlab 1, radiology | 4 | 3.5 | 52 | 0 | Normal | machine shifting to b2; leakage test failed; tidal volume not getting | NEEDS PARTS/REPLACEMENT: flow sensor replaced and flow sensor acalibration done | LOW
141 | ENDSC/ECG/01 | Ecg Machine | GASTRO ICU, GMICU | 4 | 3.2 | 38 | 0 | Normal | no battery backup; not turning on; battery drained | changed the battery and rectified the issue | LOW
142 | OT/CO2L/01 | Co2 Laser | Main OT | 3 | 4.1 | 82 | 0 | Normal | scanner mirror problem; mirror checking | NEEDS PARTS/REPLACEMENT: checked and need to place the mirror properly. informed engg tomorrow will come; serviced with engg and now ok. need to replace mirror temporary adjusted | LOW
143 | NSICU/VEN/03 | Ventilator | MDICU, NSICU | 4 | 3 | 30 | 0 | Normal | pm due; circuit check fail; handle loose | checked and cleaned the expiratory cassete,now working fine | LOW
144 | HP/IFM/01 | Immuno Fluorescence Microscope | HistoPathology, Pathology, histopathology | 4 | 2.7 | 40 | 0 | Normal | pm due; lamp expired; lamp replacement | NEEDS PARTS/REPLACEMENT: replaced the lamp; replaced lamp | LOW
145 | NSICU/VEN/02 | Ventilator | NSICU, Nsicu, nsicu | 4 | 2.6 | 26 | 0 | Normal | pm due; wheel jam; wheel loose | checked and found ok | LOW
146 | NSICU/DFB/01 | Defibrilator | NSICU, Nsicu, nsicu | 5 | 1.5 | 15 | 0 | Normal | to insert the paper; paper completed; to insert printing roll | NEEDS PARTS/REPLACEMENT: paper replaced; inserted new printing paper and found ok | LOW
147 | OT/HOL/02 | Holmium Laser | OT, Ot recovery, endoscopy | 3 | 3.6 | 31 | 0 | Normal | shifting; fiber checking; shifting to endoscopy | shifted the machine | LOW
148 | — | Sealing Machine | CSSD, cat lab, cath lab, cathlab | 4 | 2.5 | 25 | 0 | Normal | belt broken; belt came off; to change belt | NEEDS PARTS/REPLACEMENT: replaced with new one; replaced the belt and working ok | LOW
149 | NICU/IFP/08 | Infusion Pump | 3rd T2, 3rd t2, 7th main, 8TH Main, 9th south | 4 | 2.3 | 28 | 0 | Normal | to set the pump; showing alarm; continueous alarm | set the infusion and connected to patient | LOW
150 | MDICU/MON/09 | Monitor | MDICU, mdicu | 4 | 2.2 | 26 | 0 | Normal | rr value not getting; ppv not getting; saturation probe not working | checked and found ok | LOW
151 | SICU/FP/04 | Feeding Pump | GSHDU, SICU | 4 | 2.2 | 34 | 0 | Normal | not getting on; battery backup issue; door not closing | NEEDS PARTS/REPLACEMENT: replaced the battery; need to service | LOW
152 | NSICU/VEN/01 | Ventilator | NSICU, Neuro ICU, nsicu | 4 | 2.2 | 26 | 0 | Critical | for pressure check; for checking; tidal volume not achieving | checked and found ok | LOW
153 | VPSL/RD/US/03 | Ultrasound Machine | 3rd Executive, 3rd executive | 3 | 2.9 | 57 | 0 | Normal | image not saving; to clean the cursor; cursor not working properly | deleted the old files. found working | LOW
154 | ORICU/DFB/01 | Defibrillator | ORICU, Oricu, Ortho ICU, Ortho OT | 4 | 1.8 | 26 | 0 | Critical | need to change paper; print not clear; to insert printing paper | changed paper and working properly | LOW
155 | NEU/EB/01 | Electric Bed | EEG, Neuro Physio, Nsicu | 4 | 1.7 | 26 | 0 | Normal | no battery backup; fullup not working | NEEDS PARTS/REPLACEMENT: checked and need to check the control box; need to replace bed for service | LOW
156 | EMR/ECG/02 | Ecg Machine | Emergency | 4 | 1.7 | 25 | 0 | Normal | avl not getting; to print the recent result; tray broken | NEEDS PARTS/REPLACEMENT: replaced the paper tray; replaced the touch screen | LOW
157 | OT/SAR/03 | Cusa | Main OT, main ot | 3 | 2.7 | 55 | 0 | Critical | irrigation not working | checked and found ok | LOW
158 | — | Powercord | gshdu, onco icu | 3 | 2.7 | 55 | 0 | normal | plugtop damaged; complaint | NEEDS PARTS/REPLACEMENT: replaced with new one; replaced plugtop | LOW
159 | — | Dozee Monitoring System | 7th floor, New 6th block, old block | 3 | 2.6 | 76 | 0 | normal | for checking the working condition; give training to 7th floor staffs | checked the working condition | LOW
160 | — | Examination Couch | Foot and ankle, Ops, podiatry | 3 | 2.6 | 52 | 0 | Normal | to receive new examination couches; head end not working | NEEDS PARTS/REPLACEMENT: need to replace control box; checked and need to replace the same | LOW
161 | NICU/IFP/07 | Infusion Pump | 10TH new, 10th New, 3rd Floor, 6th Main, 6th new | 4 | 1.5 | 18 | 0 | Normal | to set the infusion; showing alarm; continues alarm | connected set properly and found ok | LOW
162 | OPTHAL/SLP/01 | Slit Lamp Chair | DR.Haris OP(Ophthalmology), Ophthalmology, Opthalmology | 3 | 2.5 | 50 | 0 | Normal | up and down motion not working; to connect new bulb | checked and rectified, informed se | LOW
163 | — | Multipara Monitor | Emergency, Endoscopy , Ortho ICU, csicu, ortho icu | 3 | 2.5 | 30 | 0 | Normal | for fixing the monitor on stand; bp not getting; monitor stand loose | NEEDS PARTS/REPLACEMENT: need to change bp module informed to se | LOW
164 | CSICU/MON/02 | Monitor | CSICU, csicu, cticu | 3 | 2.5 | 30 | 0 | Normal | ecg not getting; to remove the monitor from stand; cvp not getting | checked and rectified the issue | LOW
165 | 8THDX/DFB/01 | Defibrillator | 6th isolation, 8th Isolation, 8th isolation | 4 | 1.4 | 21 | 0 | Critical | battery test fail; not getting print; opertaional check failure | NEEDS PARTS/REPLACEMENT: replaced the battery with new one | LOW
166 | AMBN/DFB/02 | Defibrillator | AMBULANCE, Ambulance, ambulance | 3 | 2.2 | 34 | 0 | Normal | battery test failed; printer out of paper; paddles error | NEEDS PARTS/REPLACEMENT: checked and need to replace the battery; replaced the battery with new one | LOW
167 | 6THF/ECG/03 | Ecg Machine | 5TH main, 6th floor, 6thf | 3 | 2.1 | 25 | 0 | normal | v2not getting; printer not working; wave not getting | NEEDS PARTS/REPLACEMENT: replaced the cable | LOW
168 | RD/US/01 | Ultrasound Machine | Emergency, Radiology | 3 | 2 | 40 | 0 | Normal | image not sending; not getting worklist; pacs issues | connected the lan cable and ping the ip address now working ok | LOW
169 | — | Shaver Cable | DROT, or ot, orot | 3 | 2 | 40 | 0 | normal | checking; shaver foot switch stuck | checked after service | LOW
170 | — | Ambal | Foot and Ankle, Gastro, endoscopy | 3 | 2 | 30 | 0 | normal | image not getting; print not getting; to connect the junction box | checked and found issue with printer | LOW
171 | SICU/EB/12 | Electric Bed | 9TH NEW, MDICU, sicu | 3 | 2 | 24 | 0 | Normal | not working properly | NEEDS PARTS/REPLACEMENT: need to replace the the remote | LOW
172 | MB/RBIS/01 | Autoimmunoanalyzer | Microbiology, microbiology | 3 | 2 | 40 | 0 | Normal | to clean the machine; to clean the vitek machine; alarm | cleaned and found ok | LOW
173 | MDICU/US/01 | Ultrasound Machine | Radiology, Ultrasound, mdicu | 3 | 2 | 40 | 0 | Normal | to rearrange and properly stow the cable, avoid damage; to assist se in preventive maintenance | rearranged and stow the cable | LOW
174 | CATH/SM/01 | Sealing Machine | Cathlab, cath lab, cathlab | 3 | 1.9 | 29 | 0 | Normal | belt broken | NEEDS PARTS/REPLACEMENT: checked and changed with new one | LOW
175 | — | Oct Machine | Ophthalmology | 3 | 1.9 | 38 | 0 | Normal | showing check ink level; oct machine demo; to attend the demo | checked and rectified | LOW
176 | RD/PXRAY/08 | Portable Xray | radiology | 3 | 1.9 | 38 | 0 | normal | pacs not getting; not connecting to detector; checking battery backup | software settings reloaded and checked working | LOW
177 | MDICU/ECG/01 | Ecg Machine | 6th south, MDICU, Mdicu | 3 | 1.8 | 22 | 0 | Normal | clamb broken; ecg artifacts; to set the time | NEEDS PARTS/REPLACEMENT: replaced with new one | LOW
178 | NICU/IFP/05 | Infusion Pump | 10th main, 3rd Extension , 6th North, 6th north | 3 | 1.8 | 26 | 0 | Normal | to set the infusion; to set the pump; showing alarm | set the infusion | LOW
179 | MB/CF/02 | Centrifuge | Microbiology, microbiology | 3 | 1.8 | 26 | 0 | Normal | not working | checked, but issue exist | LOW
180 | 9THF/EB/01 | Electric Bed | 2nd old, Cardiac stepdown HDU, physiotherapy | 3 | 1.7 | 34 | 0 | Normal | electric bed not working; foot end stuck; 3pin plug top complaint | checked and rectified | LOW
181 | RD/CR/01 | Classic Dr | Emergency, radiology, xray | 3 | 1.6 | 32 | 0 | Normal | keyboard not working; cr reader shifting; image not erasing properly | restarted the system and found working | LOW
182 | MB/BSC/02 | Bio Safety Cabinet | microbiology | 3 | 1.5 | 45 | 0 | normal | flow sensor cable loose; sealing work assistance; to check all the connection | service done by sr engg | LOW
183 | — | Dozee Vital Monitor | 3rd floor, 4th floor | 3 | 1.5 | 30 | 0 | normal | for checking the availibility of rooms to do electrical work | checked and repoerted to electrical department and ip coordinators | LOW
184 | — | Video Laryngoscope | emergency | 3 | 1.5 | 30 | 0 | normal | demo; low battery; demo of video laryngoscope | demo given and handed over the machine | LOW
185 | CSSD/ETO/02 | Eto Machine | CSSD | 3 | 1.4 | 27 | 0 | Normal | fan error; error | checked and rectified issue | LOW
186 | NEPH/DM/22 | Dialysis Machine | Dialysis, dialysis | 3 | 1.4 | 28 | 0 | Normal | pm due; high venous pressure | pm to be carried out | LOW
187 | EMR/MON/12 | Multipara Monitor | Emergency, Emergency Department, emergency | 3 | 1.4 | 21 | 0 | normal | etco2 not getting; to tight the monitor stand; nibp not monitoring | NEEDS PARTS/REPLACEMENT: replaced with new one | LOW
188 | HP/SAM/02 | Microtome | Histopathology, histopathology | 3 | 1.3 | 27 | 0 | Normal | need to reinstall the microtome; blade holder stuck; section thick and thin | re installed the same | LOW
189 | SICU/MON/24 | Multiparamonitor | SICU | 3 | 1.3 | 27 | 0 | Normal | monitor loose from wall stand; nibp not monitoring; arterial bp not monitoring | tightened the monitor and found ok | LOW
190 | ORICU/MON/10 | Multipara Monitor | Blood bank, Ortho ICU | 3 | 1.2 | 23 | 0 | Normal | to remove monitor from the stand; to set the auto interval of nibp; to set the monitor | NEEDS PARTS/REPLACEMENT: removed and replaced to other stand | LOW
191 | CSICU/MON/05 | Monitor | CSICU | 3 | 1.2 | 17 | 0 | Normal | to remove the monitor from stand; to fix the monitor; fix on stand | removed the monitor from stand | LOW
192 | EMR/MON/13 | Multipara Monitor | Emergency, emergency | 3 | 1.1 | 22 | 0 | Non critical | to fix the monitor; nibp not monitoring; need to calibrate spo2 | fixed the monitor properly , found ok | LOW
193 | EMR/MON/14 | Multipara Monitor | Emergency, emergency | 3 | 1.1 | 22 | 0 | Non critical | nibp not monitoring; ecg calibration; spo2 calibrate | checked and found okay kept under observation | LOW
194 | HP/SAM/04 | Semi Auto Microtome | Histopathology, Pathology | 3 | 1 | 30 | 0 | Normal | blade loose; allignment isue | service done by sr engr | LOW
195 | RD/DR/01 | Dr | Radiology | 3 | 0.8 | 23 | 0 | Normal | image not sending to pacs; image not sending; collimator light not working | NEEDS PARTS/REPLACEMENT: replaced the bulb and working ok | LOW
196 | RAD/PDR/07 | Portable Dr | radiology | 3 | 0.7 | 14 | 0 | normal | wifi not getting; detector not get connect to wifi | checked and rectified | LOW

## Biomed vs Breakdown
Biomed Purchase Request vs. Breakdown Analysis ¦ FWD Call Register Apr 2025 – Mar 2026
Biomed Request | Equipment in Call Register | Total Calls | Total Downtime (hrs) | Avg Downtime/ Call (min) | Breakdown Flag | Verdict | Analysis
TMT Machine | Tmt M/C (row 112) | 3 | 10.6 | 212 | LOW | NOT VALIDATED | Low call volume, zero pending calls. Issues are mechanical (belt slip, incline failure). Clinical expansion case, not breakdown-driven.
Injector (Cath Lab) | Dye Injector (row 116) | 3 | 9 | 270 | LOW | NOT VALIDATED | Only PM-related calls. High avg downtime from a single event. Weak breakdown case — purchase may be capacity-driven.
Thomson Refractor (Liver Surgery) | Not in register | — | — | — | — | CANNOT VALIDATE | No breakdown history in call register. New clinical requirement, not a replacement purchase.
Ultrasound Machine (Radiology & Medical Van) | Multiple units (rows 45, 48, 63, 70, 72, 88–89, 102, 109, 152) | 5–12 per unit | Up to 19.5 | 80–167 | HIGH / MEDIUM | VALIDATED | Multiple units with persistent PACS connectivity, probe, and display failures. Radiology units flagged HIGH. Strong case.
Infusion Pump | Infusion Pump (rows 20, 95, 148, 160, 177) | 24 | 20 | 41 | HIGH | VALIDATED | 24 calls across multiple departments, recurring alarm and setting failures, 9% pending. Supports replacement/augmentation.
Syringe Pump | Syringe Pump (row 21) | 22 | 21.3 | 64 | HIGH | VALIDATED | 22 calls, battery and mechanical failures across several ICUs, 9% pending. Supports purchase.
ICU Pendants (14-bed MDICU) | Pendant (row 87) | 9 | 14.6 | 73 | MEDIUM | PARTIALLY VALIDATED | Existing pendants show suction and PM issues (MEDIUM flag). New MDICU is a greenfield build — purchase is expansion-driven, but breakdown history confirms category reliability risk.
Trauma OT Table | Ot Table (row 43) | 14 | 19.5 | 65 | HIGH | VALIDATED | 14 calls, remote and tilt failures, avg 65 min downtime per incident. Clear replacement case.
Note: "Calls" and downtime figures for equipment without a unique ID are fleet-wide totals across all departments. For unit-level replacement decisions, filter by Equipment ID in the Equipment Breakdown sheet.

## Summary
Equipment Junk/Replace Analysis — Summary
Category | Count | Avg Calls | Avg Downtime (hrs)
HIGH — Replace | 60 | 14 | 16.4
MEDIUM — Monitor | 45 | 7 | 7
LOW — Maintain | 91 | 4 | 3.5
Source: 13 monthly call registers, Apr 2025 – Mar 2026